Supplementary Table 1. Comparison of respirator approval standards for KF masks and EU FFP respirators

| Variables | Filtering efficiency (%) | Test agent | Inhalation resistance- pressure drop (flow rate) | Total inward leakagea (%) |
| --- | --- | --- | --- | --- |
| Korea |
| KF80 | 80 | NaCl | ≤ 60 (at 30 L/min) | 25 |
| KF94 | 94 | NaCl & paraffin oil | ≤ 70 (at 30 L/min) | 11 |
| KF99 | 99 | NaCl & paraffin oil | ≤ 100 (at 30 L/min) | 5 |
| Europe |
| FFP1 | 80 | NaCl & paraffin oil | ≤ 60 (at 30 L/min), ≤ 210 (at 95 L/min) | 22b |
| FFP2 | 94 | NaCl & paraffin oil | ≤ 70 (at 30 L/min), ≤ 240 (at 95 L/min) | 8b |
| FFP3 | 99 | NaCl & paraffin oil | ≤ 100 (at 30 L/min), ≤ 300 (at 95 L/min) | 2b |

FFP = filtering facepiece.

aAt least 46 of the 50 individual exercise results (10 subjects × 5 exercises) for total inward leakage shall not be greater than the requirements.

bFor European standards, at least 8 out of 10 individual wearers' arithmetic means for total inward leakage shall not be greater than the requirements as well.