Supplementary Table 1. SARS-CoV-2 results and Ct values of the three assays and STANDARD M nCoV

| Specimen No. | Results and Ct values if they are positive |  |  |  |  |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  | STANDARD M nCoV (Ct cut-off 36) |  |  | PowerChek (Ct cut-off 38) |  |  | STANDARD M (Ct cut-off 36) |  |  | Allplex (Ct cut-off: N and S 38, RdRp 40) |  |  |  |
|  | Result | E | ORF1ab | Result | E | ORF1ab | Result | E | ORF1ab | Result | N | RdRp | S |
| 1 | P | 27.1 | 25.8 | P | 29.1 | 31.1 | P | 28.3 | 27.0 | P | 28.7 | 27.6 | 27.8 |
| 2 | P | 33.0 | 32.7 | P | 35.1 | 37.2 | P | 32.8 | 32.4 | P | 34.0 | 33.3 | 34.4 |
| 3 | P | 24.2 | 22.9 | P | 25.6 | 28.5 | P | 30.5 | 29.3 | P | 27.5 | 25.8 | 25.3 |
| 4 | P | 33.9 | 32.3 | P | 35.0 | 37.1 | P | 24.6 | 23.8 | P | 35.2 | 33.6 | 33.5 |
| 5 | P | 29.8 | 28.9 | P | 32.1 | 34.4 | P | 33.9 | 31.9 | P | 31.6 | 30.6 | 30.5 |
| 6 | P | 32.6 | 31.0 | P | 33.8 | 37.5 | P | 34.4 | 30.5 | I | - | 34.7 | 32.8 |
| 7 | P | 29.0 | 27.6 | P | 30.8 | 32.8 | P | 29.4 | 27.8 | P | 29.6 | 30.2 | 29.1 |
| 8 | P | 31.2 | 30.4 | P | 33.9 | 36.0 | P | 32.9 | 31.2 | P | 33.2 | 32.5 | 32.3 |
| 9 | P | 31.0 | 29.4 | P | 33.1 | 35.0 | P | 32.1 | 32.3 | P | 33.0 | 31.2 | 32.0 |
| 10 | P | 25.7 | 24.1 | P | 27.7 | 29.9 | P | 26.6 | 25.0 | P | 27.4 | 26.5 | 26.0 |
| 11 | P | 29.8 | 28.4 | P | 31.6 | 33.5 | P | 29.9 | 28.2 | P | 31.4 | 30.4 | 29.9 |
| 12 | P | 28.2 | 27.1 | P | 30.2 | 32.2 | P | 29.1 | 27.9 | P | 29.9 | 29.5 | 28.7 |
| 13 | P | 25.2 | 23.8 | P | 27.2 | 29.1 | P | 26.8 | 25.5 | P | 26.5 | 25.8 | 25.3 |
| 14 | P | 32.4 | 30.5 | P | 33.3 | 36.4 | P | 33.1 | 31.7 | P | 34.3 | 33.7 | 32.9 |
| 15 | P | 28.8 | 27.5 | P | 30.1 | 32.5 | P | 28.4 | 27.0 | P | 30.1 | 29.3 | 29.2 |
| 16 | P | 32.2 | 31.3 | P | 33.7 | 35.8 | P | 32.7 | 29.5 | P | 33.5 | 33.7 | 32.8 |
| 17 | P | 29.4 | 28.6 | P | 31.9 | 34.0 | P | 30.1 | 28.7 | P | 29.8 | 30.4 | 30.4 |
| 18 | P | 31.2 | 30.0 | P | 32.7 | 35.0 | P | 31.1 | 30.1 | P | 33.0 | 31.8 | 32.2 |
| 19 | P | 31.6 | 30.2 | P | 32.5 | 34.6 | P | 30.8 | 28.7 | P | 31.2 | 31.3 | 31.3 |
| 20 | P | 32.0 | 31.0 | P | 34.7 | 35.6 | P | 32.9 | 29.6 | P | 33.8 | 32.5 | 34.2 |
| 21 | P | 23.3 | 22.0 | P | 25.1 | 26.7 | P | 22.4 | 20.7 | P | 24.1 | 23.4 | 23.5 |
| 22 | P | 24.8 | 23.4 | P | 26.8 | 29.1 | P | 25.4 | 24.1 | P | 25.6 | 25.1 | 25.3 |
| 23 | P | 33.1 | 31.0 | P | 33.8 | 34.7 | P | 35.3 | 30.5 | P | 33.6 | 32.8 | 33.6 |
| 24 | P | 32.7 | 32.0 | P | 35.5 | 35.6 | P | 31.1 | 30.3 | P | 35.0 | 33.4 | 33.5 |
| 25 | P | 26.1 | 25.3 | P | 28.0 | 30.7 | P | 26.2 | 25.8 | P | 26.6 | 27.3 | 26.5 |
| 26 | P | 29.7 | 28.5 | P | 31.6 | 33.6 | P | 29.7 | 28.4 | P | 31.2 | 30.6 | 30.6 |
| 27 | P | 24.3 | 22.8 | P | 26.3 | 28.4 | P | 25.1 | 23.3 | P | 25.9 | 24.7 | 25.1 |
| 28 | P | 29.4 | 27.6 | P | 31.3 | 33.2 | P | 29.9 | 27.9 | P | 29.8 | 29.6 | 29.4 |
| 29 | P | 31.8 | 31.3 | P | 34.7 | 36.7 | P | 30.7 | 29.5 | P | 33.0 | 33.2 | 32.7 |
| 30 | P | 23.9 | 22.5 | P | 25.8 | 27.9 | P | 24.8 | 23.4 | P | 24.9 | 24.4 | 24.4 |
| 31 | P | 32.7 | 31.4 | P | 34.5 | 36.1 | P | 35.1 | 31.7 | P | 33.4 | 33.5 | 33.2 |
| 32 | P | 26.1 | 24.5 | P | 27.8 | 30.2 | P | 26.6 | 25.0 | P | 27.2 | 26.4 | 26.3 |
| 33 | P | 24.3 | 23.1 | P | 26.1 | 28.2 | P | 24.6 | 23.1 | P | 25.5 | 24.7 | 24.1 |
| 34 | P | 29.5 | 28.0 | P | 31.1 | 33.3 | P | 29.7 | 27.9 | P | 29.9 | 29.6 | 29.8 |
| 35 | P | 29.5 | 29.3 | P | 31.7 | 34.0 | P | 29.6 | 28.9 | P | 30.3 | 31.1 | 30.6 |
| 36 | P | 34.7 | 33.2 | I | 35.8 | - | P | 34.0 | 33.1 | I | - | 35.3 | 34.0 |
| 37 | P | 31.8 | 32.3 | P | 35.1 | 36.3 | P | 32.7 | 32.7 | P | 34.6 | 34.8 | 33.6 |
| 38 | P | 26.2 | 24.6 | P | 28.2 | 30.3 | P | 28.5 | 27.3 | P | 28.0 | 27.1 | 26.9 |
| 39 | P | 27.3 | 26.2 | P | 29.4 | 32.0 | P | 28.0 | 26.8 | P | 29.1 | 28.4 | 27.8 |
| 40 | P | 25.4 | 24.1 | P | 27.4 | 30.1 | P | 26.9 | 25.5 | P | 26.9 | 25.5 | 25.7 |
| 41 | P | 27.4 | 26.3 | P | 29.4 | 32.1 | P | 28.2 | 26.8 | P | 29.0 | 29.2 | 28.7 |
| 42 | P | 26.3 | 24.3 | P | 28.3 | 30.4 | P | 30.0 | 28.1 | P | 28.0 | 26.6 | 26.8 |
| 43 | P | 30.1 | 28.6 | P | 32.0 | 34.6 | P | 31.1 | 29.5 | P | 30.4 | 31.1 | 30.6 |
| 44 | P | 28.3 | 27.2 | P | 30.6 | 32.8 | P | 30.0 | 28.4 | P | 30.1 | 29.5 | 29.0 |
| 45 | P | 24.4 | 23.0 | P | 26.5 | 29.1 | P | 25.2 | 23.5 | P | 26.8 | 25.6 | 25.3 |
| 46 | P | 18.3 | 16.5 | P | 20.3 | 21.5 | P | 19.0 | 16.8 | P | 19.8 | 18.5 | 18.9 |
| 47 | P | 31.3 | 30.3 | P | 33.8 | 37.7 | P | 33.5 | 30.9 | P | 34.3 | 33.7 | 32.0 |
| 48 | P | 31.7 | 30.3 | P | 33.8 | 35.6 | P | 31.7 | 31.5 | P | 35.0 | 33.4 | 33.2 |
| 49 | P | 29.1 | 27.9 | P | 31.5 | 33.4 | P | 30.1 | 29.1 | P | 30.7 | 29.6 | 29.3 |
| 50 | P | 32.1 | 30.8 | P | 34.8 | 35.6 | P | 33.0 | 31.3 | P | 32.8 | 34.3 | 32.5 |
| 51 | P | 31.2 | 30.2 | P | 33.0 | 35.8 | P | 31.8 | 30.2 | P | 32.4 | 31.3 | 32.1 |
| 52 | P | 28.9 | 27.6 | P | 31.1 | 33.2 | P | 29.5 | 28.3 | P | 30.9 | 29.4 | 29.6 |
| 53 | P | 31.7 | 30.1 | P | 33.1 | 35.2 | P | 34.0 | 30.6 | P | 32.2 | 31.8 | 32.3 |
| 54 | P | 25.8 | 24.3 | P | 28.1 | 30.3 | P | 26.8 | 25.5 | P | 27.6 | 26.6 | 26.6 |
| 55 | P | 32.2 | 30.8 | P | 35.3 | 36.4 | P | 35.0 | 31.8 | P | 35.0 | 33.1 | 32.8 |
| 56 | P | 31.4 | 30.1 | P | 32.5 | 35.5 | P | 32.9 | 33.3 | P | 34.6 | 32.2 | 32.0 |
| 57 | P | 32.3 | 32.5 | P | 35.6 | 37.6 | P | 35.6 | 31.9 | P | 32.6 | 33.8 | 36.7 |
| 58 | P | 28.2 | 26.4 | P | 30.5 | 32.4 | P | 30.6 | 29.1 | P | 29.4 | 27.8 | 28.7 |
| 59 | P | 31.9 | 30.2 | P | 34.3 | 35.4 | P | 31.8 | 31.1 | P | 32.7 | 32.7 | 31.9 |
| 60 | P | 26.9 | 25.2 | P | 29.0 | 30.7 | P | 28.1 | 26.5 | P | 28.3 | 27.2 | 27.2 |
| 61 | P | 30.2 | 28.8 | P | 31.6 | 33.5 | P | 32.8 | 30.0 | P | 33.7 | 31.3 | 31.4 |
| 62 | P | 25.4 | 24.0 | P | 27.6 | 30.2 | P | 25.5 | 24.0 | P | 27.3 | 26.4 | 25.5 |
| 63 | P | 30.2 | 29.1 | P | 32.9 | 34.4 | P | 31.2 | 29.8 | P | 33.1 | 30.7 | 30.9 |
| 64 | P | 24.6 | 23.0 | P | 26.4 | 28.9 | P | 26.5 | 24.8 | P | 26.1 | 25.2 | 24.3 |
| 65 | P | 34.6 | 33.1 | I | 36.7 | - | I | 36.4 | 35.1 | P | 35.2 | 35.7 | 36.0 |
| 66 | P | 31.3 | 30.0 | I | 38.3 | 35.1 | N | - | 36.4 | P | 34.9 | 31.7 | 32.3 |
| 67 | P | 33.1 | 32.2 | P | 36.0 | 37.6 | I | - | 33.5 | P | 35.7 | 34.8 | 34.9 |
| 68 | P | 24.2 | 23.0 | P | 26.4 | 28.5 | P | 27.1 | 25.6 | P | 26.1 | 25.4 | 25.3 |
| 69 | P | 29.0 | 28.0 | P | 31.0 | 33.2 | P | 31.5 | 29.6 | P | 31.2 | 30.3 | 29.6 |
| 70 | P | 32.1 | 31.0 | P | 34.8 | 36.8 | P | 32.9 | 33.0 | P | 33.7 | 33.6 | 34.3 |
| 71 | P | 22.5 | 21.2 | P | 24.7 | 27.1 | P | 24.1 | 22.4 | P | 24.0 | 23.4 | 23.3 |
| 72 | P | 35.4 | 32.1 | I | - | 37.2 | P | 35.3 | 33.4 | P | 34.7 | 36.8 | 36.0 |
| 73 | P | 28.3 | 27.1 | P | 30.7 | 32.6 | P | 30.6 | 28.4 | P | 30.6 | 29.0 | 28.5 |
| 74 | P | 34.1 | 32.2 | P | 34.5 | 37.6 | P | 33.4 | 35.0 | P | 35.4 | 35.8 | 34.6 |
| 75 | P | 32.9 | 31.3 | P | 36.8 | 36.7 | I | - | 33.1 | I | 35.2 | 36.2 | - |
| 76 | P | 22.0 | 20.3 | P | 24.0 | 26.0 | P | 23.7 | 21.9 | P | 24.2 | 22.6 | 22.7 |
| 77 | P | 26.9 | 25.4 | P | 29.3 | 31.5 | P | 31.0 | 28.9 | P | 28.8 | 27.7 | 27.7 |
| 78 | P | 22.0 | 20.5 | P | 24.3 | 26.2 | P | 24.0 | 22.2 | P | 24.3 | 22.9 | 22.8 |
| 79 | P | 23.9 | 22.0 | P | 26.0 | 27.2 | P | 24.9 | 23.2 | P | 25.4 | 24.1 | 24.4 |
| 80 | P | 31.1 | 29.6 | N | - | - | P | 33.9 | 31.0 | P | 34.1 | 32.5 | 31.4 |
| 81 | P | 25.6 | 24.4 | P | 28.0 | 30.2 | P | 27.3 | 25.6 | P | 27.8 | 26.2 | 26.4 |
| 82 | P | 28.0 | 27.1 | P | 29.8 | 32.0 | P | 30.1 | 28.9 | P | 29.2 | 29.8 | 28.8 |
| 83 | P | 18.1 | 16.6 | P | 20.3 | 22.0 | P | 20.3 | 19.4 | P | 19.7 | 18.7 | 18.8 |
| 84 | P | 25.2 | 23.8 | P | 27.4 | 29.8 | P | 27.5 | 26.1 | P | 27.0 | 25.8 | 26.0 |
| 85 | P | 27.2 | 25.8 | P | 29.5 | 32.1 | P | 28.8 | 27.6 | P | 29.7 | 27.7 | 28.0 |
| 86 | P | 29.2 | 28.2 | P | 31.2 | 33.4 | P | 30.2 | 29.1 | P | 30.7 | 30.4 | 29.9 |
| 87 | P | 25.4 | 24.3 | P | 27.9 | 30.1 | P | 27.3 | 26.5 | P | 27.2 | 26.6 | 25.6 |
| 88 | P | 20.9 | 19.3 | P | 23.1 | 24.3 | P | 23.9 | 22.4 | P | 21.8 | 21.4 | 21.3 |
| 89 | P | 20.7 | 19.1 | P | 23.1 | 24.3 | P | 22.9 | 21.1 | P | 22.4 | 21.3 | 21.1 |
| 90 | P | 18.7 | 17.2 | P | 20.9 | 22.3 | P | 21.0 | 19.5 | P | 20.0 | 19.2 | 19.6 |
| 91 | P | 29.0 | 27.5 | P | 31.1 | 33.2 | P | 30.5 | 29.2 | P | 32.6 | 30.2 | 29.8 |
| 92 | P | 30.0 | 29.1 | P | 32.4 | 34.3 | P | 31.9 | 30.7 | P | 31.4 | 30.8 | 30.9 |
| 93 | P | 26.6 | 25.5 | P | 28.8 | 31.2 | P | 28.6 | 28.0 | P | 28.0 | 27.0 | 26.9 |
| 94 | P | 32.2 | 30.9 | I | 35.1 | 38.1 | P | 33.3 | 32.6 | P | 33.4 | 33.6 | 32.5 |
| 95 | P | 24.6 | 23.3 | P | 26.9 | 29.2 | P | 27.0 | 25.8 | P | 26.4 | 25.4 | 24.6 |
| 96 | P | 27.7 | 26.9 | P | 30.2 | 32.1 | P | 33.4 | 32.4 | P | 29.3 | 28.4 | 28.8 |
| 97 | P | 34.3 | 33.3 | N | 38.7 | 38.3 | P | 34.8 | 34.7 | I | - | 35.4 | 36.1 |
| 98–298 | N | - | - | N | - | - | N | - | - | N | - | - | - |

SARS-CoV-2 = severe acute respiratory syndrome-coronavirus-2, Ct = cycle threshold, P = positive, I = inconclusive, N = negative.