Supplementary Table 3. Discordant samples among AdvanSure and three kits regarding influenza A/B and RSV

| Case No. | Result, target, and Ct values | RT-PCR and gene sequencing results |
| --- | --- | --- |
|  | AdvanSure (Ct cut-off 27) | PowerChek (Ct cut-off 38) | STANDARD M (Ct cut-off 36) | Allplex (Ct cut-off 38) |  |
|  | Result | Ct value | Result | Ct value | Result | Ct value | Result | Ct value |  |
| Case 1 | Flu A | Flu A: 26.9 | Negative | Flu A: none | Flu A | Flu A: 34.0 | Negative | Flu A: none | Not tested |
| Case 2 | Flu A | RSV: none | Flu A |  | Flu A |  | Flu A, RSV | RSV: 36.7 | No amplification of RSV gene |
| Case 3 | Flu A | RSV: none | Flu A |  | Flu A |  | Flu A, RSV | RSV: 33.0 | RSV positive |
| Case 4 | Flu B | Flu A: none | Flu B | Flu A: none | Flu B | Flu A: none | Flu A, Flu B | Flu A: 34.0 | Not tested |
| Case 5 | Rhinovirus A/B/C | RSV: none | Negative |  | Negative |  | RSV | RSV: 35.9 | RSV positive |
| Case 6 | Influenza B, rhinovirus A/B/C | RSV: none | Flu B |  | Flu B |  | Flu B, RSV | RSV: 34.9 | RSV positive |
| Case 7 | Influenza B, adenovirus, rhinovirus A/B/C | RSV: none | Flu B |  | Flu B |  | Flu B, RSV | RSV: 34.4 | RSV positive |

RSV = respiratory syncytial virus, Ct = cycle threshold, RT-PCR = reverse transcription polymerase chain reaction, Flu = influenza virus.